Standard Operating Procedure - Measurements of Ultraviolet
Fluorescence During Oil Spills

Sampling
Water is run from approximately the middle of a Niskin bottle and collected as a 20 ml aliquot into a standard scintillation vial with a foil-lined cap. The sample is then acidified with 20 l of 10% HCl and stored at 4oC in the dark. Three-ml aliquots of the sample can then be analyzed by scanning fluorometry ashore or by ship-board, fixed-wavelength fluorometry (for rapid feedback).

Scanning Fluorescence Analysis (Shore Based)
Samples are transferred into plastic (methacrylate) or UV-grade cuvettes for fluorescence measurements in a Shimadzu RF5301 scanning fluorometer located onshore. Each sample is subjected to ultraviolet (UV) irradiation with the excitation energy fixed at 280 nm. Emission energy is then tracked and recorded as the fluorometer scans between 300 and 545 nm.
The emission spectra are stored as Panorama fluorescence files and exported to an Excel workbook to plot the spectra and calculate/plot the fluorescence intensity ratio (FIR) for each spectrum.

Fixed Wavelength Fluorescence Analysis (Ship Based)
Two fixed wavelength UV fluorometers (Quantech/Thermo Scientific) are employed in tandem to determine fluorescence intensity ratios (FIRs). One fluorometer is equipped with a pair of wavelength filters allowing excitation at 280 nm and emission at 340 nm. The second fluorometer is equipped with the same 280nm excitation filter and a longer (445 nm) wavelength.
Two individual aliquots of 3 ml are transferred to either methacrylate or UV grade quartz cuvettes, and emission at 340 nm and 445 nm recorded on the respective fluorometers. FIRs can then be calculated from fluorescence intensity at 340 nm divided by intensity at 445 nm.