LISST-100X Particle Size Analyzer Standard Operation Procedure

General Description

The Sequoia Laser In-situ Scattering and Transmissometry (LISST-100X) instrument uses the technique of laser diffraction to determine particle size-distribution (PSD) and/or volume distribution. It consists of an optical system that produces a collimated laser beam, a detector array, electronics for signal pre-amplification and processing, a data storage and scheduling computer, and a battery system.

List of Materials

- User’s Manual with software disk
- LISST-100X instrument
- Bench top instrument stands
- Communication cable
- External power cable
- Small volume horizontal test chamber
- Allen wrench set
- Pre-moistened alcohol wipes
- De-ionized water

Bench-top standard operation procedures

1. Remove the instrument stands and set them up on a flat bench top working surface.
2. Remove the LISST-100X from the case and set it on the stands.
3. Check the optical windows. Clean the Transmit window and the Receive window by wiping them with a pre-moistened wipe, and then by rinsing them with de-ionized water.
4. Remove the small volume horizontal test chamber from the case. Attach the flexible tubing and tubing stop clamp.
5. Slip the chamber between the optical windows of the instrument such that the optic can be submerged for calibration and testing (see Appendix F of the User’s manual for more detail of chamber installation)
6. Check and adjust the chamber and the spacer to make sure that the chamber attachment to the Transmit and the Receiver windows are sealed properly without leaking.
7. Fill the small chamber with clean de-ionized water for calibration
8. Remove the Communication cable for the case. Connect the 5-pin underwater connector of the cable to the LISST-100X, and the 9-pin DB-9 of the cable to a lap-top personal computer.
9. Install software for communication and processing program, including the calibration files specific for the instrument. Create a shortcut of the LISST-100X program on desktop for the ease of operation.
10. Start the LISST-SOP software by clicking on the LISST.exe file.
11. Select Wake up LISST from the LISST menu. A dialogue box will appear counting down the maximum time (138 seconds) required for the instrument to wake up.
12. Acquire background scattering profiles by selecting Collect Background Scatter Data from the LISST menu or by pressing the button on the toolbar. The background scattering measurement is critical to good instrument performance and will also check the overall health of the instrument. It will verify that all of the systems are functioning and that the optics in proper alignment. The current background will be acquired and displayed relative to the factory background scattering for the instrument (specific to each unit).
13. Press the BEGIN Collection button to acquire 20 samples and the display the average background scattering after the collection. If the values are acceptable and the values can be saved to a file, enter the filename in the box and press the Accept and Save button. If there is a problem, an error message will be displayed. Take corrective actions and make sure the current laser power is close to the Factory laser power.
14. Choose Open Real-Time Session from the File menu or press the button to open the Real-Time session. Choose a background file to use when processing the data. Choose an output PSD file. A display similar to this will appear,

The left hand figure displays a bar chart showing the volume concentration in each of the 32 log spaced size classes. The right hand plot will be the Cumulative Concentration. To view the samples as a movie press the Timer Off button. The slider bar next to the button adjusts the refresh rate. The First, Prev, Next, and Last buttons allow you to step through the measurements one frame at a time. When the last frame is reached the movie display will stop. The slider bar controls the data acquisition time. The Scale +/- Button adjusts the Particle Size Distribution scale.
1. After finishing with the instrument, press LISST to sleep from the LISST menu to put it back into its low power sleep mode.
2. Choose Convert binary PSD to ASCII file under the File menu, save file as .xls file for data processing and manipulation in Excel.

Vessel side in-water continuous monitoring standard operation procedures

1. Connect the LISST-100X to the PC and start the LISST-SOP program.
2. Open the Operating Modes window by choosing Operating Modes from the LISST menu or by pressing the button on the toolbar. Select the Operating Mode Tab at the top of the window. Choose Fixed Sample Rate. Select Start Conditions tab to configure when the instrument will begin sampling among five options: Depth, Time, External Mechanical Switch, External Digital Input, and Time Delay. Select the mode by clicking on the button next to the its label. Select the correct parameters as required.
3. Choose the Stop Conditions Tab to select the conditions when sampling should stop. The available stop conditions are: Depth, Time. External Mechanical Switch, External Digital Input, Fixed number of samples, and Maximum Memory or Low Battery.
4. Select Apply or OK buttons to configure the instrument with the current settings. If the Apply button is pressed the program will return to the current window. Returning to the Instrument Status window will display a summary of the current settings. If the OK button is pressed, when the configuration is complete the user will be prompted to open the Terminal window to start the program. To start the program and have it start looking for the Start conditions press the Run button on the Terminal window.
5. Once the program is confirmed to be running and waiting for the correct start conditions, the LISST-SOP program can be closed and the communications cable can be disconnected.
6. Replace the connector cap before deployment.
7. Deploy the configured LISST in water beside the vessel to carry out in situ monitoring of in situ particle size distribution in the water column.
8. When the start conditions are met and the instrument will start sampling. The green LED on the connector endcap will illuminate each time a sample is acquired. The instrument will continue collecting data until the pre-set stop conditions are met.
9. Stop a running program by pressing the Stop button on the Terminal window or the button.
10. Press the Instrument Query Button to display the instrument status including the number of samples saved.
11. Select Offload from the LISST menu or choose the button from the toolbar. A list of files will appear. Choose the files to offload by clicking on them while hold down the CTRL key. The data is offloaded at 115K baud.
12. To process the data file choose Open Raw Data file or press the button on the toolbar. Select the instrument serial number, raw data file to open, and background scatter file to use when processing the data. Press the Process File button to convert the raw file into processed size distributions.
13. Processed data files are stored as binary files with a PSD extension. The PSD files can be opened and viewed on the screen. To open a processed Particle Size Distribution file choose Open Particle Distribution File or press the button on the toolbar. The steps are the same as bench top operation.