## OV149
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
OV149 | SW-20100801-OV12-001 | 1 |  | 5.33 | 5.71 |  | 5.68 | 5.76 | 1149
OV149 | SW-20100801-OV12-002 | 2 |  | 5.23 | 5.97 |  | 6.04 | 5.81 | 1010
OV149 | SW-20100801-OV12-007 | 5 |  | 3.38 | 3.88 |  | 4.2 | 3.91 | 400
OV149 | SW-20100801-OV12-010 | 7 |  | 5.48 | 7.2 |  | 7.56 | 7.34 | 50

## OV150
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
OV150 | SW-20100801-OV12-013 | 1 |  | 4.57 | 4.9 |  | 5.03 | 4.94 | 750
OV150 | SW-20100801-OV12-014 | 2 |  | 3.58 | 4.36 |  | 4.62 | 4.39 | 650
OV150 | SW-20100801-OV12-015 | 3 |  | 3.42 | 3.93 |  | 4.29 | 3.97 | 500
OV150 | SW-20100801-OV12-016 | 4 |  | 3.39 | 4.13 |  | 4.65 | 4.19 | 300
OV150 | SW-20100801-OV12-017 | 5 |  | 3.25 | 3.46 |  | 4.37 | 4.03 | 233
OV150 | SW-20100801-OV12-018 | 6 |  | 3.81 | 4.89 |  | 5.3 | 4.97 | 128
OV150 | SW-20100801-OV12-019 | 7 |  | 5.51 | 6.08 |  | 7.45 | 7.24 | 50
OV150 | SW-20100801-OV12-020 | 8 |  | 4.83 | 5.32 |  | 6.75 | 6.44 | 3

## OV151
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
OV151 | SW-20100801-OV12-021 | 1 |  | 5.22 | 5.73 |  | 5.77 | 5.77 | 1060
OV151 | SW-20100801-OV12-022 | 2 |  | 4.91 | 5.63 |  | 5.89 | 5.68 | 1020
OV151 | SW-20100801-OV12-023 | 3 |  | 3.52 | 4.22 |  | 4.56 | 4.2 | 600
OV151 | SW-20100801-OV12-024 | 4 |  | 3.53 | 3.38 |  | 4.28 | 3.92 | 351
OV151 | SW-20100801-OV12-025 | 5 |  | 3.46 | 4.21 |  | 4.57 | 4.24 | 300
OV151 | SW-20100801-OV12-026 | 6 |  | 3.91 | 4.83 |  | 5.24 | 4.89 | 150
OV151 | SW-20100801-OV12-029 | 7 |  | 5.14 | 7.17 |  | 7.61 | 7.43 | 50
OV151 | SW-20100801-OV12-031 | 8 |  | 4.29 | 6.35 |  | 6.95 | 6.51 | 3

## OV152
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
OV152 | SW-20100801-OV12-032 | 1 |  | 6.12 | 5.87 |  | 6.78 | 6.5 | 1350
OV152 | SW-20100801-OV12-033 | 2 |  | 5.73 | 6.06 |  | 6.76 | 6.51 | 1300
OV152 | SW-20100801-OV12-034 | 3 |  | 5.84 | 6.46 |  | 6.99 | 6.49 | 1250
OV152 | SW-20100801-OV12-035 | 4 |  | 5.59 | 6.26 |  | 6.71 | 5.92 | 1192
OV152 | SW-20100801-OV12-036 | 5 |  | 5.38 | 6.1 |  | 6.31 | 5.83 | 1165
OV152 | SW-20100801-OV12-037 | 6 |  | 5.73 | 7.45 |  | 7.73 | 3.98 | 400
OV152 | SW-20100801-OV12-038 | 7 |  | 6.03 | 6.8 |  | 7.64 | 7.29 | 50
OV152 | SW-20100801-OV12-039 | 8 |  | 4.38 | 4.66 |  | 6.89 | 6.45 | 3

## OV153
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
OV153 | SW-20100802-OV12-001 | 1 |  | 6.25 | 6.79 |  | 6.61 | 6.78 | 1450
OV153 | SW-20100802-OV12-004 | 2 |  | 5.86 | 5.72 |  | 6.67 | 6.39 | 1300
OV153 | SW-20100802-OV12-005 | 3 |  | 5.44 | 5.87 |  | 6.3 | 5.69 | 1206
OV153 | SW-20100802-OV12-012 | 4 |  |  | 5.83 |  |  | 5.72 | 1157
OV153 | SW-20100802-OV12-007 | 5 |  | 4.61 | 5.36 |  | 5.89 | 5.75 | 1050
OV153 | SW-20100802-OV12-009 | 6 |  |  | 3.81 |  |  | 3.78 | 400
OV153 | SW-20100802-OV12-010 | 7 |  | 5.35 | 6.87 |  | 7.4 | 7.17 | 50
OV153 | SW-20100802-OV12-011 | 8 |  | 4.58 | 6.35 |  | 6.89 | 6.49 | 3

## OV154
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
OV154 | SW-20100802-OV12-013 | 1 |  | 5.63 | 6.07 |  | 7.02 | 6.63 | 1350
OV154 | SW-20100802-OV12-014 | 2 |  | 5.64 | 6.31 |  | 6.82 | 6.29 | 1272
OV154 | SW-20100802-OV12-015 | 3 |  | 5.15 | 5.67 |  | 6.38 | 5.92 | 1205
OV154 | SW-20100802-OV12-016 | 4 |  | 5.34 | 5.84 |  | 6.42 | 5.59 | 1169
OV154 | SW-20100802-OV12-018 | 6 |  | 3.15 | 3.8 |  | 4.26 | 3.84 | 400
OV154 | SW-20100802-OV12-019 | 7 |  | 5.26 | 6.73 |  | 7.34 | 7.11 | 50
OV154 | SW-20100802-OV12-020 | 8 |  | 4.22 | 6.32 |  | 6.86 | 6.48 | 3

## OV155
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
OV155 | SW-20100802-OV12-021 | 1 |  | 5.05 | 5.87 |  | 6.08 | 5.91 | 1060
OV155 | SW-20100802-OV12-022 | 2 |  | 5.18 | 5.83 |  | 6.22 | 5.84 | 1020

## OV156
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
OV156 | SW-20100802-OV12-032 | 1 |  |  | 6.69 |  | 6.26 | 6.75 | 1340
OV156 | SW-20100802-OV12-033 | 2 |  |  | 6.54 |  | 6.66 | 6.56 | 1250
OV156 | SW-20100802-OV12-034 | 3 |  |  | 6.35 |  | 6.7 | 6.41 | 1200
OV156 | SW-20100802-OV12-035 | 4 |  |  | 5.33 |  | 5.76 | 6.16 | 1150
OV156 | SW-20100802-OV12-036 | 5 |  |  | 5.84 |  | 6.33 | 4.73 | 1131
OV156 | SW-20100802-OV12-039 | 6 |  |  | 4.55 |  | 5.01 | 4.63 | 1124
OV156 | SW-20100802-OV12-040 | 7 |  |  | 5.13 |  | 5.52 | 5.1 | 1100
OV156 | SW-20100802-OV12-041 | 8 |  |  | 5.36 |  | 5.84 | 5.38 | 1085
OV156 | SW-20100802-OV12-042 | 9 |  |  | 5.81 |  | 6.29 | 5.83 | 1025

## OV157
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
OV157 | SW-20100802-OV12-048 | 3 |  |  | 6.79 |  | 7.36 | 6.83 | 1500
OV157 | SW-20100802-OV12-049 | 4 |  |  | 6.87 |  | 7.25 | 6.77 | 1400
OV157 | SW-20100802-OV12-050 | 5 |  |  | 6.62 |  | 7.17 | 6.66 | 1300
OV157 | SW-20100802-OV12-051 | 6 |  |  | 6.09 |  | 6.65 | 6.43 | 1200
OV157 | SW-20100802-OV12-052 | 7 |  |  | 5.61 |  | 6.18 | 5.46 | 1132
OV157 | SW-20100802-OV12-053 | 8 |  |  | 5.86 |  | 6.42 | 5.86 | 1069

## BM141
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM141 | SW-20100729-BM15-7 | 7 |  | 5.65 |  | 5.3922 | 6.63 | 6.34 | 2
BM141 | SW-20100729-BM15-5 | 5 |  | 5.82 |  | 4 | 6.9 | 3.91 | 500
BM141 | SW-20100729-BM15-3 | 3 |  | 6.02 |  | 5.3861 | 7.09 | 6.25 | 1150
BM141 | SW-20100729-BM15-2 | 2 |  | 6.11 |  | 5.95 | 7.1 | 6.42 | 1200
BM141 | SW-20100729-BM15-1 | 1 |  | 6.05 |  | 6.2393 | 7.4 | 6.74 | 1500

## BM142
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM142 | SW-20100729-BM15-13 | 6 |  | 4.7 |  | 4.9314 | 6.25 | 4.79 | 250
BM142 | SW-20100729-BM15-12 | 5 |  | 3.95 |  | 4.1662 | 5.35 | 3.85 | 500
BM142 | SW-20100729-BM15-10 | 3 |  | 5.62 |  | 6.12 | 6.95 | 6.32 | 1150
BM142 | SW-20100729-BM15-9 | 2 |  | 5.8 |  | 6.0975 | 6.98 | 6.44 | 1200
BM142 | SW-20100729-BM15-8 | 1 |  | 6 |  | 6.2451 | 7.17 | 6.9 | 1500

## BM143
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM143 | SW-20100729-BM15-21 | 6 |  | 4.45 |  | 4.22 | 5.38 | 4.29 | 250
BM143 | SW-20100729-BM15-20 | 5 |  | 4.05 |  | 3.9482 | 5.01 | 3.84 | 500
BM143 | SW-20100729-BM15-18 | 3 |  | 5.88 |  | 5.9704 | 6.95 | 6.31 | 1150
BM143 | SW-20100729-BM15-17 | 2 |  | 5.86 |  | 6 | 7.04 | 6.41 | 1200
BM143 | SW-20100729-BM15-16 | 1 |  | 6.07 |  | 6.06 | 7.25 | 6.77 | 1500

## BM144
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM144 | SW-20100729-BM15-29 | 6 |  | 4.68 |  | 4.8104 | 5.85 | 4.72 | 250
BM144 | SW-20100729-BM15-28 | 5 |  | 5.5 |  | 4 | 6.75 | 3.95 | 500
BM144 | SW-20100729-BM15-26 | 3 |  | 5.8 |  | 6.0588 | 7.01 | 6.28 | 1150
BM144 | SW-20100729-BM15-25 | 2 |  | 5.95 |  | 6.2 | 7.06 | 6.34 | 1200
BM144 | SW-20100729-BM15-24 | 1 |  | 6.2 |  | 6.42 | 7.2 | 6.71 | 1500

## BM145
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM145 | SW-20100730-BM15-6 | 6 |  | 5.74 |  | 4.2 | 5.65 | 4.53 | 250
BM145 | SW-20100730-BM15-5 | 5 |  | 5.18 |  | 3.56 | 5.27 | 3.92 | 500
BM145 | SW-20100730-BM15-3 | 3 |  | 7.12 |  | 5.9821 | 7.1 | 6.29 | 1150
BM145 | SW-20100730-BM15-2 | 2 |  | 7.03 |  | 6.08 | 7.12 | 6.41 | 1200
BM145 | SW-20100730-BM15-1 | 1 |  | 7.67 |  | 6.2 | 7.42 | 6.88 | 1500

## BM146
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM146 | SW-20100730-BM15-14 | 6 |  | 5.37 |  | 4.243 | 5.48 | 4.77 | 250
BM146 | SW-20100730-BM15-13 | 5 |  | 4.9 |  | 3.8508 | 4.92 | 3.88 | 500
BM146 | SW-20100730-BM15-11 | 3 |  | 7.12 |  | 5.68 | 6.94 | 6.33 | 1150
BM146 | SW-20100730-BM15-10 | 2 |  | 7.23 |  | 6.0339 | 7.01 | 6.47 | 1200
BM146 | SW-20100730-BM15-9 | 1 |  | 7.5 |  | 6.0539 | 7.2 | 6.81 | 1500

## BM147
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM147 | SW-20100730-BM15-24 | 6 |  | 5.62 |  | 4.5074 | 5.58 | 4.62 | 250
BM147 | SW-20100730-BM15-23 | 5 |  |  |  | 5.6 |  | 5.92 | 1050
BM147 | SW-20100730-BM15-21 | 3 |  | 6.41 |  | 5.7059 | 6.8 | 6.11 | 1150
BM147 | SW-20100730-BM15-20 | 2 |  | 6.65 |  | 6.3546 | 7.1 | 6.29 | 1200
BM147 | SW-20100730-BM15-19 | 1 |  | 6.6 |  | 6.0827 | 7.2 | 6.62 | 1500

## BM148
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM148 | SW-20100730-BM15-33 | 6 |  | 5.14 |  | 4.08 | 5.58 | 4.63 | 250
BM148 | SW-20100730-BM15-32 | 5 |  | 4.85 |  | 3.6731 | 5.33 | 3.89 | 500
BM148 | SW-20100730-BM15-30 | 3 |  | 6.35 |  | 6 | 6.98 | 6.31 | 1150
BM148 | SW-20100730-BM15-29 | 2 |  | 6.58 |  | 5.9478 | 7.15 | 6.43 | 1200
BM148 | SW-20100730-BM15-28 | 1 |  | 6.79 |  | 6 | 7.28 | 6.56 | 1500

## BM149
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM149 | SW-20100731-BM15-6 | 9 |  | 5.65 |  | 4.55 | 5.84 | 4.69 | 250
BM149 | SW-20100731-BM15-4 | 7 |  | 6.17 |  | 5.38 | 6.47 | 5.82 | 1050
BM149 | SW-20100731-BM15-3 | 5 |  | 6.26 |  | 5.13 | 6.45 | 5.59 | 1150
BM149 | SW-20100731-BM15-2 | 3 |  | 6.38 |  | 5.52 | 6.52 | 5.65 | 1190
BM149 | SW-20100731-BM15-1 | 1 |  | 6.75 |  | 5.54 | 6.92 | 6.22 | 1300

## BM154
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM154 | SW-20100804BM16-1 | 2 |  | 6.28 | 6.03 | 6.3 | 6.3264 | 5.97 | 1268
BM154 | SW-20100804BM16-2 | 4 |  | 5.52 | 5.116 | 5.38 | 5.5336 | 4.97 | 1215
BM154 | SW-20100804BM16-3 | 6 |  | 6.39 | 5.893 | 6.28 | 6.2455 | 5.84 | 1195
BM154 | SW-20100804BM16-4 | 8 |  | 6.54 | 6.041 | 6.39 | 6.3587 | 6.03 | 1174
BM154 | SW-20100804BM16-5 | 10 |  | 4.66 | 4.163 | 4.27 | 4.5628 | 3.93 | 500

## BM150
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM150 | SW-20100731-BM15-13 | 7 |  | 6.67 |  | 5.34 | 6.6 | 5.74 | 1050
BM150 | SW-20100731-BM15-12 | 5 |  | 6.18 |  | 5.27 | 6.52 | 5.45 | 1100
BM150 | SW-20100731-BM15-11 | 4 |  | 7.11 |  | 6.19 | 7.14 | 6.27 | 1200
BM150 | SW-20100731-BM15-10 | 3 |  | 6.74 |  | 5.05 | 6.71 | 5.71 | 1245
BM150 | SW-20100731-BM15-9 | 1 |  | 7.11 |  | 5.99 | 7.09 | 6.61 | 1450

## BM151
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM151 | SW-20100731-BM15-23 | 9 |  | 5.03 |  | 4.29 | 5.48 | 4.53 | 250
BM151 | SW-20100731-BM15-22 | 8 |  | 5.04 |  | 3.83 | 5.12 | 3.97 | 500
BM151 | SW-20100731-BM15-20 | 5 |  | 6.2 |  | 5.69 | 6.4 | 5.7 | 1125
BM151 | SW-20100731-BM15-19 | 3 |  | 6.41 |  | 5.5 | 6.48 | 5.73 | 1150
BM151 | SW-20100731-BM15-18 | 1 |  | 6.7 |  | 5.87 | 6.83 | 5.97 | 1195

## BM152
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM152 | SW-20100731-BM15-33 | 10 |  | 4.6 |  | 4.27 | 5.1 | 4.45 | 250
BM152 | SW-20100731-BM15-30 | 8 |  | 5.01 |  | 3.89 | 5.49 | 3.88 | 500
BM152 | SW-20100731-BM15-29 | 6 |  | 5.3 |  | 4.66 | 5.99 | 4.8 | 750
BM152 | SW-20100731-BM15-28 | 4 |  | 6.15 |  | 5.05 | 6.59 | 5.54 | 950
BM152 | SW-20100731-BM15-27 | 2 |  | 6.7 |  | 5.66 | 6.87 | 5.84 | 1050

## BM153
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM153 | SW-20100731-BM15-42 | 6 |  | 4.43 |  | 4.04 | 5.24 | 4.17 | 250
BM153 | SW-20100731-BM15-41 | 5 |  | 4.64 |  | 4.05 | 5.02 | 3.88 | 500
BM153 | SW-20100731-BM15-40 | 4 |  | 5.18 |  | 4.4 | 5.71 | 4.76 | 750
BM153 | SW-20100731-BM15-39 | 3 |  | 5.96 |  | 6.05 | 6.7 | 6.06 | 1050
BM153 | SW-20100731-BM15-38 | 2 |  | 5.91 |  | 5.73 | 6.94 | 6.38 | 1150

## BM155
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM155 | SW-20100804BM16-9 | 2 |  | 5.62 | 6.322 | 6.04 | 6.67 | 6.34 | 1274
BM155 | SW-20100804BM16-10 | 4 |  | 5.49 | 4.096 | 5.92 | 5.99 | 5.51 | 1212
BM155 | SW-20100804BM16-11 | 6 |  | 5.16 | 4.427 | 4.91 | 5.65 | 5.27 | 1201
BM155 | SW-20100804BM16-12 | 8 |  | 5.71 | 4.737 | 5.89 | 6.24 | 5.93 | 1174
BM155 | SW-20100804BM16-13 | 10 |  | 4.15 | 3.841 | 3.81 | 4.5 | 3.92 | 500

## BM156
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM156 | SW-20100804BM16-17 | 2 |  | 5.56 | 6.429 | 6.08 | 6.57 | 6.34 | 1274
BM156 | SW-20100804BM16-18 | 4 |  | 5.33 | 5.231 | 6.16 | 6.56 | 6.24 | 1219
BM156 | SW-20100804BM16-19 | 6 |  | 5.52 | 4.54 | 5.62 | 6.41 | 6.11 | 1214
BM156 | SW-20100804BM16-20 | 8 |  | 4.81 | 4.954 | 5.04 | 5.5 | 4.97 | 1178
BM156 | SW-20100804BM16-21 | 10 |  | 4.01 | 3.723 | 3.87 | 4.76 | 3.92 | 500

## BM157
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM157 | SW-20100804BM16-24 | 2 |  | 5.8 | 6.485 | 6.24 | 6.88 | 6.7 | 1449
BM157 | SW-20100804BM16-25 | 4 |  | 5.23 | 4.876 | 5.54 | 6.47 | 6.07 | 1219
BM157 | SW-20100804BM16-26 | 6 |  | 5.52 | 5.327 | 5.64 | 6.29 | 5.83 | 1194
BM157 | SW-20100804BM16-27 | 8 |  | 5.33 | 5.819 | 5.75 | 6.05 | 5.72 | 1159
BM157 | SW-20100804BM16-28 | 10 |  | 4.25 | 3.781 | 3.92 | 4.75 | 3.94 | 500

## BM158
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM158 | SW-20100805BM16-1 | 2 |  | 6.06 |  | 6.99 | 7.16 | 6.92 | 1684
BM158 | SW-20100805BM16-2 | 4 |  | 5.22 | 5.556 | 6.03 | 6.54 | 5.92 | 1249
BM158 | SW-20100805BM16-3 | 6 |  | 5.15 | 5.716 | 5.58 | 6.13 | 5.62 | 1199
BM158 | SW-20100805BM16-4 | 8 |  | 5.23 | 5.341 | 5.66 | 6.08 | 5.67 | 1149
BM158 | SW-20100805BM16-5 | 10 |  | 3.78 | 3.738 | 3.89 | 4.79 | 3.92 | 500

## BM159
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM159 | SW-20100805BM16-9 | 2 |  | 5.89 | 5.372 | 6.25 | 6.89 | 6.7 | 1402
BM159 | SW-20100805BM16-10 | 4 |  | 5.34 | 5.354 | 5.75 | 6.22 | 5.67 | 1222
BM159 | SW-20100805BM16-11 | 6 |  | 5.1 | 4.288 | 5.12 | 6.04 | 5.59 | 1181
BM159 | SW-20100805BM16-12 | 8 |  | 4.93 | 5.679 | 5.44 | 5.97 | 5.76 | 1153
BM159 | SW-20100805BM16-13 | 10 |  | 3.86 | 3.704 | 3.85 | 4.68 | 3.93 | 500

## BM160
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM160 | SW-20100805BM16-16 | 2 |  | 6.16 | 6.122 | 6.61 | 6.86 | 6.75 | 1449
BM160 | SW-20100805BM16-17 | 4 |  | 5.36 | 4.924 | 5.4 | 6.21 | 5.91 | 1184
BM160 | SW-20100805BM16-18 | 6 |  | 5.25 | 4.942 | 5.29 | 6.1 | 5.72 | 1147
BM160 | SW-20100805BM16-19 | 8 |  | 4.96 | 4.932 | 4.82 | 5.45 | 4.88 | 749
BM160 | SW-20100805BM16-20 | 10 |  | 3.83 | 4.586 | 3.92 | 4.73 | 3.95 | 500

## BM161
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM161 | SW-20100805BM16-23 | 2 |  | 5.78 | 5.38 | 6.34 | 6.96 | 6.77 | 1449
BM161 | SW-20100805BM16-24 | 4 |  | 5.43 | 4.854 | 6.01 | 6.58 | 6.2 | 1275
BM161 | SW-20100805BM16-25 | 6 |  | 5.58 | 4.833 | 5.61 | 5.91 | 5.79 | 1140
BM161 | SW-20100805BM16-26 | 8 |  | 4.53 | 4.34 | 4.59 | 5.33 | 4.83 | 749
BM161 | SW-20100805BM16-27 | 10 |  | 3.83 | 2.858 | 3.84 | 4.63 | 3.94 | 500

## BM162
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM162 | SW-20100806BM16-2 | 2 |  | 5.02 | 5.336 | 5.35 | 5.66 | 5.42 | 949
BM162 | SW-20100806BM16-3 | 3 |  | 4.77 |  | 4.85 | 5.49 | 5.29 | 899
BM162 | SW-20100806BM16-4 | 4 |  | 4.3 | 4.881 | 4.97 | 5.31 | 5.07 | 850
BM162 | SW-20100806BM16-5 | 5 |  | 4.32 | 4.35 | 4.6 | 5.11 | 4.86 | 800
BM162 | SW-20100806BM16-6 | 6 |  | 4.25 | 4.557 | 4.47 | 4.98 | 4.75 | 750
BM162 | SW-20100806BM16-7 | 7 |  | 4.45 | 4.65 | 4.56 | 4.92 | 4.68 | 700
BM162 | SW-20100806BM16-8 | 8 |  | 4.1 | 4.142 | 4.13 | 4.61 | 4.34 | 649
BM162 | SW-20100806BM16-9 | 9 |  | 3.65 | 3.9 | 4.02 | 4.43 | 4.14 | 599
BM162 | SW-20100806BM16-10 | 10 |  | 3.86 | 3.949 | 3.89 | 4.24 | 3.96 | 550
BM162 | SW-20100806BM16-11 | 11 |  | 3.48 | 3.619 | 3.79 | 4.19 | 3.86 | 500

## BM167
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM167 | SW-20100806BM16-1 | 1 |  | 5.22 | 5.275 | 5.82 | 6.13 | 5.84 | 1079
BM167 | SW-20100806BM16-9 | 9 |  | 3.86 | 3.577 | 3.77 | 4.79 | 3.88 | 500

## BM169
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM169 | SW-20100810BM17-1 | 1 |  | 7.07 | 6.693 | 6.57 | 6.76 | 6.69 | 1299
BM169 | SW-20100810BM17-2 | 2 |  | 6.46 | 6.146 | 6.03 | 6.64 | 6.38 | 1199
BM169 | SW-20100810BM17-3 | 3 |  | 6.53 | 5.919 | 6.31 | 6.51 | 6.35 | 1149
BM169 | SW-20100810BM17-4 | 4 |  | 6.46 | 6.044 | 5.92 | 6.47 | 6.22 | 1139
BM169 | SW-20100810BM17-5 | 5 |  | 5.82 | 6.064 | 5.93 | 6.28 | 6.01 | 1127
BM169 | SW-20100810BM17-6 | 6 |  | 5.18 | 5.722 | 5.61 | 6.04 | 5.74 | 1120
BM169 | SW-20100810BM17-7 | 7 |  | 5.75 | 5.518 | 5.51 | 6.04 | 5.74 | 1109
BM169 | SW-20100810BM17-8 | 8 |  | 5.3 | 5.604 | 5.66 | 6.01 | 5.71 | 1099
BM169 | SW-20100810BM17-9 | 9 |  | 5.8 | 5.817 | 5.77 | 6.08 | 5.8 | 1089
BM169 | SW-20100810BM17-10 | 10 |  | 6.07 | 5.597 | 5.6 | 6.12 | 5.86 | 1079
BM169 | SW-20100810BM17-11 | 11 |  | 5.78 | 6.327 | 5.99 | 6.17 | 6.01 | 1069
[5 empty] | 0.82831 | 0.93283

## BM170
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM170 | SW-20100810BM17-13 | 1 |  | 6.51 | 6.706 | 6.36 | 7.13 | 6.98 | 1798
BM170 | SW-20100810BM17-14 | 2 |  | 6.31 | 6.959 | 6.41 | 7.1 | 6.96 | 1599
BM170 | SW-20100810BM17-15 | 3 |  | 6.45 | 6.676 | 6.45 | 6.79 | 6.88 | 1399
BM170 | SW-20100810BM17-16 | 4 |  | 6.46 | 6.742 | 6.53 | 6.78 | 6.62 | 1199
BM170 | SW-20100810BM17-17 | 5 |  | 5.82 | 6.052 | 5.73 | 6.43 | 6.19 | 1024
BM170 | SW-20100810BM17-18 | 6 |  | 5.6 | 6.112 | 5.6 | 6.29 | 6.1 | 999
BM170 | SW-20100810BM17-19 | 7 |  | 5.4 | 5.146 | 5.05 | 5.54 | 5.27 | 799
BM170 | SW-20100810BM17-20 | 8 |  | 4.3 | 4.437 | 4.22 | 4.71 | 4.39 | 600
BM170 | SW-20100810BM17-21 | 9 |  | 3.98 | 3.995 | 3.87 | 4.35 | 3.98 | 400
BM170 | SW-20100810BM17-22 | 10 |  | 4.19 | 3.925 | 4.21 | 4.78 | 4.51 | 200
BM170 | SW-20100810BM17-23 | 11 |  | 6.15 | 5.301 | 6.19 | 6.41 | 6.38 | 2
[5 empty] | 0.95179 | 0.9889

## BM171
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM171 | SW-20100810BM17-25 | 1 |  | 6.32 | 6.3 | 6.79 | 7.17 | 7 | 1689
BM171 | SW-20100810BM17-26 | 2 |  | 6.05 | 6.416 | 6.9 | 7.08 | 6.98 | 1599
BM171 | SW-20100810BM17-27 | 3 |  | 6.62 | 6.407 | 6.79 | 7.1 | 6.89 | 1399
BM171 | SW-20100810BM17-28 | 4 |  | 6.13 | 5.35 | 5.67 | 6.52 | 6.68 | 1199
BM171 | SW-20100810BM17-29 | 5 |  | 6.01 | 5.73 | 5.62 | 6.27 | 6.08 | 1000
BM171 | SW-20100810BM17-30 | 6 |  | 5.25 | 5.075 | 5.32 | 5.61 | 5.34 | 800
BM171 | SW-20100810BM17-31 | 7 |  | 4.25 | 4.284 | 4.19 | 4.75 | 4.39 | 600
BM171 | SW-20100810BM17-32 | 8 |  | 4.42 | 3.943 | 3.74 | 4.31 | 3.94 | 400
BM171 | SW-20100810BM17-33 | 9 |  | 4.15 | 4.023 | 4.24 | 4.76 | 4.48 | 200
BM171 | SW-20100810BM17-34 | 10 |  | 5.65 | 5.54 | 5.51 | 6.52 | 6.41 | 2
[5 empty] | 0.95628 | 0.95676

## BM172
StationID | SampleID | BottleNumber | DO_LaMotte | Membrane | Winkler | Hach | Optical | Seabird | Depth
BM172 | SW-20100811BM17-1 | 1 |  |  | 6.112 | 6.18 | 6.78 | 6.63 | 1343 | 0.45
BM172 | SW-20100811BM17-2 | 2 |  |  | 5.677 | 6.05 | 6.54 | 6.36 | 1199 | 0.31
BM172 | SW-20100811BM17-3 | 3 |  |  | 5.591 | 6.06 | 6.32 | 6 | 1189 | -0.06
BM172 | SW-20100811BM17-4 | 4 |  |  | 5.042 | 5.63 | 6.16 | 5.84 | 1179 | 0.21
BM172 | SW-20100811BM17-5 | 5 |  |  | 4.997 | 6 | 6.19 | 6.02 | 1169 | 0.02
BM172 | SW-20100811BM17-6 | 6 |  |  | 4.813 | 5.67 | 6.06 | 5.88 | 1159 | 0.21
BM172 | SW-20100811BM17-7 | 7 |  |  | 4.46 | 5.58 | 5.98 | 5.79 | 1149 | 0.21
BM172 | SW-20100811BM17-8 | 8 |  |  | 4.87 | 5.79 | 5.97 | 5.73 | 1139 | -0.06
BM172 | SW-20100811BM17-9 | 9 |  |  | 4.655 | 5.36 | 5.91 | 5.71 | 1129 | 0.35
BM172 | SW-20100811BM17-10 | 10 |  |  | 4.824 | 5.37 | 5.91 | 5.73 | 1119 | 0.36
BM172 | SW-20100811BM17-11 | 11 |  |  | 4.787 | 5.81 | 5.93 | 5.75 | 1110 | -0.06
[5 empty] | 0.90971 | 0.78848

## BM_DATA
StationID | SampleID | Niskin Bottle# | DO_Membrane | DO_Winkler | DO_Hach | DO_Optical | DO_Seabird | Depth
|  |  | (mg/L) | (mg/L) | (mg/L) | (mg/L) | (mg/L) | Depth [m]
BM141 | SW-20100729-BM15-7 | 7 | 5.65 |  | 5.3922 | 6.63 | 6.34 | 2
BM141 | SW-20100729-BM15-5 | 5 | 5.82 |  | 4 | 6.9 | 3.91 | 500
BM141 | SW-20100729-BM15-3 | 3 | 6.02 |  | 5.3861 | 7.09 | 6.25 | 1150
BM141 | SW-20100729-BM15-2 | 2 | 6.11 |  | 5.95 | 7.1 | 6.42 | 1200
BM141 | SW-20100729-BM15-1 | 1 | 6.05 |  | 6.2393 | 7.4 | 6.74 | 1500
BM142 | SW-20100729-BM15-13 | 6 | 4.7 |  | 4.9314 | 6.25 | 4.79 | 250
BM142 | SW-20100729-BM15-12 | 5 | 3.95 |  | 4.1662 | 5.35 | 3.85 | 500
BM142 | SW-20100729-BM15-10 | 3 | 5.62 |  | 6.12 | 6.95 | 6.32 | 1150
BM142 | SW-20100729-BM15-9 | 2 | 5.8 |  | 6.0975 | 6.98 | 6.44 | 1200
BM142 | SW-20100729-BM15-8 | 1 | 6 |  | 6.2451 | 7.17 | 6.9 | 1500
BM143 | SW-20100729-BM15-21 | 6 | 4.45 |  | 4.22 | 5.38 | 4.29 | 250
BM143 | SW-20100729-BM15-20 | 5 | 4.05 |  | 3.9482 | 5.01 | 3.84 | 500
BM143 | SW-20100729-BM15-18 | 3 | 5.88 |  | 5.9704 | 6.95 | 6.31 | 1150
BM143 | SW-20100729-BM15-17 | 2 | 5.86 |  | 6 | 7.04 | 6.41 | 1200
BM143 | SW-20100729-BM15-16 | 1 | 6.07 |  | 6.06 | 7.25 | 6.77 | 1500
BM144 | SW-20100729-BM15-29 | 6 | 4.68 |  | 4.8104 | 5.85 | 4.72 | 250
BM144 | SW-20100729-BM15-28 | 5 | 5.5 |  | 4 | 6.75 | 3.95 | 500
BM144 | SW-20100729-BM15-26 | 3 | 5.8 |  | 6.0588 | 7.01 | 6.28 | 1150
BM144 | SW-20100729-BM15-25 | 2 | 5.95 |  | 6.2 | 7.06 | 6.34 | 1200
BM144 | SW-20100729-BM15-24 | 1 | 6.2 |  | 6.42 | 7.2 | 6.71 | 1500
BM145 | SW-20100730-BM15-6 | 6 | 5.74 |  | 4.2 | 5.65 | 4.53 | 250
BM145 | SW-20100730-BM15-5 | 5 | 5.18 |  | 3.56 | 5.27 | 3.92 | 500
BM145 | SW-20100730-BM15-3 | 3 | 7.12 |  | 5.9821 | 7.1 | 6.29 | 1150
BM145 | SW-20100730-BM15-2 | 2 | 7.03 |  | 6.08 | 7.12 | 6.41 | 1200
BM145 | SW-20100730-BM15-1 | 1 | 7.67 |  | 6.2 | 7.42 | 6.88 | 1500
BM146 | SW-20100730-BM15-14 | 6 | 5.37 |  | 4.243 | 5.48 | 4.77 | 250
BM146 | SW-20100730-BM15-13 | 5 | 4.9 |  | 3.8508 | 4.92 | 3.88 | 500
BM146 | SW-20100730-BM15-11 | 3 | 7.12 |  | 5.68 | 6.94 | 6.33 | 1150
BM146 | SW-20100730-BM15-10 | 2 | 7.23 |  | 6.0339 | 7.01 | 6.47 | 1200
BM146 | SW-20100730-BM15-9 | 1 | 7.5 |  | 6.0539 | 7.2 | 6.81 | 1500
BM147 | SW-20100730-BM15-24 | 6 | 5.62 |  | 4.5074 | 5.58 | 4.62 | 250
BM147 | SW-20100730-BM15-23 | 5 |  |  | 5.6 |  | 5.92 | 1050
BM147 | SW-20100730-BM15-21 | 3 | 6.41 |  | 5.7059 | 6.8 | 6.11 | 1150
BM147 | SW-20100730-BM15-20 | 2 | 6.65 |  | 6.3546 | 7.1 | 6.29 | 1200
BM147 | SW-20100730-BM15-19 | 1 | 6.6 |  | 6.0827 | 7.2 | 6.62 | 1500
BM148 | SW-20100730-BM15-33 | 6 | 5.14 |  | 4.08 | 5.58 | 4.63 | 250
BM148 | SW-20100730-BM15-32 | 5 | 4.85 |  | 3.6731 | 5.33 | 3.89 | 500
BM148 | SW-20100730-BM15-30 | 3 | 6.35 |  | 6 | 6.98 | 6.31 | 1150
BM148 | SW-20100730-BM15-29 | 2 | 6.58 |  | 5.9478 | 7.15 | 6.43 | 1200
BM148 | SW-20100730-BM15-28 | 1 | 6.79 |  | 6 | 7.28 | 6.56 | 1500
BM149 | SW-20100731-BM15-6 | 9 | 5.65 |  | 4.55 | 5.84 | 4.69 | 250
BM149 | SW-20100731-BM15-4 | 7 | 6.17 |  | 5.38 | 6.47 | 5.82 | 1050
BM149 | SW-20100731-BM15-3 | 5 | 6.26 |  | 5.13 | 6.45 | 5.59 | 1150
BM149 | SW-20100731-BM15-2 | 3 | 6.38 |  | 5.52 | 6.52 | 5.65 | 1190
BM149 | SW-20100731-BM15-1 | 1 | 6.75 |  | 5.54 | 6.92 | 6.22 | 1300
BM150 | SW-20100731-BM15-13 | 7 | 6.67 |  | 5.34 | 6.6 | 5.74 | 1050
BM150 | SW-20100731-BM15-12 | 5 | 6.18 |  | 5.27 | 6.52 | 5.45 | 1100
BM150 | SW-20100731-BM15-11 | 4 | 7.11 |  | 6.19 | 7.14 | 6.27 | 1200
BM150 | SW-20100731-BM15-10 | 3 | 6.74 |  | 5.05 | 6.71 | 5.71 | 1245
BM150 | SW-20100731-BM15-9 | 1 | 7.11 |  | 5.99 | 7.09 | 6.61 | 1450
BM151 | SW-20100731-BM15-23 | 9 | 5.03 |  | 4.29 | 5.48 | 4.53 | 250
BM151 | SW-20100731-BM15-22 | 8 | 5.04 |  | 3.83 | 5.12 | 3.97 | 500
BM151 | SW-20100731-BM15-20 | 5 | 6.2 |  | 5.69 | 6.4 | 5.7 | 1125
BM151 | SW-20100731-BM15-19 | 3 | 6.41 |  | 5.5 | 6.48 | 5.73 | 1150
BM151 | SW-20100731-BM15-18 | 1 | 6.7 |  | 5.87 | 6.83 | 5.97 | 1195
BM152 | SW-20100731-BM15-33 | 10 | 4.6 |  | 4.27 | 5.1 | 4.45 | 250
BM152 | SW-20100731-BM15-30 | 8 | 5.01 |  | 3.89 | 5.49 | 3.88 | 500
BM152 | SW-20100731-BM15-29 | 6 | 5.3 |  | 4.66 | 5.99 | 4.8 | 750
BM152 | SW-20100731-BM15-28 | 4 | 6.15 |  | 5.05 | 6.59 | 5.54 | 950
BM152 | SW-20100731-BM15-27 | 2 | 6.7 |  | 5.66 | 6.87 | 5.84 | 1050 |  |  |  |  | Optical probe values recorded in situ in the Niskins on deck and then in the lab (as usually measured) to provide an indication of differences caused by drawing off the samples and transporting them across the deck
BM153 | SW-20100731-BM15-42 | 6 | 4.43 |  | 4.04 | 5.24 | 4.17 | 250 | [5 empty] | LAB | Niskin
BM153 | SW-20100731-BM15-41 | 5 | 4.64 |  | 4.05 | 5.02 | 3.88 | 500 | [5 empty] | BM161 optical lab | BM161 optical
BM153 | SW-20100731-BM15-40 | 4 | 5.18 |  | 4.4 | 5.71 | 4.76 | 750 |  |  |  |  | DO | 6.96 | 6.72 | -0.24
BM153 | SW-20100731-BM15-39 | 3 | 5.96 |  | 6.05 | 6.7 | 6.06 | 1050 | [5 empty] | 6.58 | 6.12 | -0.46
BM153 | SW-20100731-BM15-38 | 2 | 5.91 |  | 5.73 | 6.94 | 6.38 | 1150 | [5 empty] | 5.91 | 5.85 | -0.06
BM154 | SW-20100804BM16-1 | 2 | 6.28 | 6.03 | 6.3 | 6.3264 | 5.97 | 1268 | [5 empty] | 5.33 | 5.2 | -0.13
BM154 | SW-20100804BM16-2 | 4 | 5.52 | 5.116 | 5.38 | 5.5336 | 4.97 | 1215 | [5 empty] | 4.63 | 4.48 | -0.15
BM154 | SW-20100804BM16-3 | 6 | 6.39 | 5.893 | 6.28 | 6.2455 | 5.84 | 1195 | [7 empty] | -0.208 | Lower DO in Niskin
BM154 | SW-20100804BM16-4 | 8 | 6.54 | 6.041 | 6.39 | 6.3587 | 6.03 | 1174
BM154 | SW-20100804BM16-5 | 10 | 4.66 | 4.163 | 4.27 | 4.5628 | 3.93 | 500 |  |  |  |  | Temp oC | 12.2 | 12.5 | 0.3
BM155 | SW-20100804BM16-9 | 2 | 5.62 | 6.322 | 6.04 | 6.67 | 6.34 | 1274 | [5 empty] | 12.5 | 12.9 | 0.4
BM155 | SW-20100804BM16-10 | 4 | 5.49 | 4.096 | 5.92 | 5.99 | 5.51 | 1212 | [5 empty] | 12.2 | 13.4 | 1.2
BM155 | SW-20100804BM16-11 | 6 | 5.16 | 4.427 | 4.91 | 5.65 | 5.27 | 1201 | [5 empty] | 12.8 | 13.6 | 0.8
BM155 | SW-20100804BM16-12 | 8 | 5.71 | 4.737 | 5.89 | 6.24 | 5.93 | 1174 | [5 empty] | 14.1 | 14.1 | 0
BM155 | SW-20100804BM16-13 | 10 | 4.15 | 3.841 | 3.81 | 4.5 | 3.92 | 500 | [7 empty] | 0.54 | Warmer temp in Niskin
BM156 | SW-20100804BM16-17 | 2 | 5.56 | 6.429 | 6.08 | 6.57 | 6.34 | 1274
BM156 | SW-20100804BM16-18 | 4 | 5.33 | 5.231 | 6.16 | 6.56 | 6.24 | 1219
BM156 | SW-20100804BM16-19 | 6 | 5.52 | 4.54 | 5.62 | 6.41 | 6.11 | 1214
BM156 | SW-20100804BM16-20 | 8 | 4.81 | 4.954 | 5.04 | 5.5 | 4.97 | 1178
BM156 | SW-20100804BM16-21 | 10 | 4.01 | 3.723 | 3.87 | 4.76 | 3.92 | 500
BM157 | SW-20100804BM16-24 | 2 | 5.8 | 6.485 | 6.24 | 6.88 | 6.7 | 1449
BM157 | SW-20100804BM16-25 | 4 | 5.23 | 4.876 | 5.54 | 6.47 | 6.07 | 1219
BM157 | SW-20100804BM16-26 | 6 | 5.52 | 5.327 | 5.64 | 6.29 | 5.83 | 1194
BM157 | SW-20100804BM16-27 | 8 | 5.33 | 5.819 | 5.75 | 6.05 | 5.72 | 1159
BM157 | SW-20100804BM16-28 | 10 | 4.25 | 3.781 | 3.92 | 4.75 | 3.94 | 500
BM158 | SW-20100805BM16-2 | 4 | 5.22 | 5.556 | 6.03 | 6.54 | 5.92 | 1249
BM158 | SW-20100805BM16-3 | 6 | 5.15 | 5.716 | 5.58 | 6.13 | 5.62 | 1199
BM158 | SW-20100805BM16-4 | 8 | 5.23 | 5.341 | 5.66 | 6.08 | 5.67 | 1149
BM158 | SW-20100805BM16-5 | 10 | 3.78 | 3.738 | 3.89 | 4.79 | 3.92 | 500
BM159 | SW-20100805BM16-9 | 2 | 5.89 | 5.372 | 6.25 | 6.89 | 6.7 | 1402
BM159 | SW-20100805BM16-10 | 4 | 5.34 | 5.354 | 5.75 | 6.22 | 5.67 | 1222
BM159 | SW-20100805BM16-11 | 6 | 5.1 | 4.288 | 5.12 | 6.04 | 5.59 | 1181
BM159 | SW-20100805BM16-12 | 8 | 4.93 | 5.679 | 5.44 | 5.97 | 5.76 | 1153
BM159 | SW-20100805BM16-13 | 10 | 3.86 | 3.704 | 3.85 | 4.68 | 3.93 | 500
BM160 | SW-20100805BM16-16 | 2 | 6.16 | 6.122 | 6.61 | 6.86 | 6.75 | 1449
BM160 | SW-20100805BM16-17 | 4 | 5.36 | 4.924 | 5.4 | 6.21 | 5.91 | 1184
BM160 | SW-20100805BM16-18 | 6 | 5.25 | 4.942 | 5.29 | 6.1 | 5.72 | 1147
BM160 | SW-20100805BM16-19 | 8 | 4.96 | 4.932 | 4.82 | 5.45 | 4.88 | 749
BM160 | SW-20100805BM16-20 | 10 | 3.83 | 4.586 | 3.92 | 4.73 | 3.95 | 500
BM161 | SW-20100805BM16-23 | 2 | 5.78 | 5.38 | 6.34 | 6.96 | 6.77 | 1449
BM161 | SW-20100805BM16-24 | 4 | 5.43 | 4.854 | 6.01 | 6.58 | 6.2 | 1275
BM161 | SW-20100805BM16-25 | 6 | 5.58 | 4.833 | 5.61 | 5.91 | 5.79 | 1140
BM161 | SW-20100805BM16-26 | 8 | 4.53 | 4.34 | 4.59 | 5.33 | 4.83 | 749
BM161 | SW-20100805BM16-27 | 10 | 3.83 | 2.858 | 3.84 | 4.63 | 3.94 | 500
BM162 | SW-20100806BM16-2 | 2 | 5.02 | 5.336 | 5.35 | 5.66 | 5.42 | 949
BM162 | SW-20100806BM16-4 | 4 | 4.3 | 4.881 | 4.97 | 5.31 | 5.07 | 850
BM162 | SW-20100806BM16-5 | 5 | 4.32 | 4.35 | 4.6 | 5.11 | 4.86 | 800
BM162 | SW-20100806BM16-6 | 6 | 4.25 | 4.557 | 4.47 | 4.98 | 4.75 | 750
BM162 | SW-20100806BM16-7 | 7 | 4.45 | 4.65 | 4.56 | 4.92 | 4.68 | 700
BM162 | SW-20100806BM16-8 | 8 | 4.1 | 4.142 | 4.13 | 4.61 | 4.34 | 649
BM162 | SW-20100806BM16-9 | 9 | 3.65 | 3.9 | 4.02 | 4.43 | 4.14 | 599
BM162 | SW-20100806BM16-10 | 10 | 3.86 | 3.949 | 3.89 | 4.24 | 3.96 | 550
BM162 | SW-20100806BM16-11 | 11 | 3.48 | 3.619 | 3.79 | 4.19 | 3.86 | 500
BM167 | SW-20100806BM16-1 | 1 | 5.22 | 5.275 | 5.82 | 6.13 | 5.84 | 1079
BM167 | SW-20100806BM16-9 | 9 | 3.86 | 3.577 | 3.77 | 4.79 | 3.88 | 500
BM169 | SW-20100810BM17-1 | 1 | 7.07 | 6.693 | 6.57 | 6.76 | 6.69 | 1299
BM169 | SW-20100810BM17-2 | 2 | 6.46 | 6.146 | 6.03 | 6.64 | 6.38 | 1199
BM169 | SW-20100810BM17-3 | 3 | 6.53 | 5.919 | 6.31 | 6.51 | 6.35 | 1149
BM169 | SW-20100810BM17-4 | 4 | 6.46 | 6.044 | 5.92 | 6.47 | 6.22 | 1139
BM169 | SW-20100810BM17-5 | 5 | 5.82 | 6.064 | 5.93 | 6.28 | 6.01 | 1127
BM169 | SW-20100810BM17-6 | 6 | 5.18 | 5.722 | 5.61 | 6.04 | 5.74 | 1120
BM169 | SW-20100810BM17-7 | 7 | 5.75 | 5.518 | 5.51 | 6.04 | 5.74 | 1109
BM169 | SW-20100810BM17-8 | 8 | 5.3 | 5.604 | 5.66 | 6.01 | 5.71 | 1099
BM169 | SW-20100810BM17-9 | 9 | 5.8 | 5.817 | 5.77 | 6.08 | 5.8 | 1089
BM169 | SW-20100810BM17-10 | 10 | 6.07 | 5.597 | 5.6 | 6.12 | 5.86 | 1079
BM169 | SW-20100810BM17-11 | 11 | 5.78 | 6.327 | 5.99 | 6.17 | 6.01 | 1069
BM170 | SW-20100810BM17-13 | 1 | 6.51 | 6.706 | 6.36 | 7.13 | 6.98 | 1798
BM170 | SW-20100810BM17-14 | 2 | 6.31 | 6.959 | 6.41 | 7.1 | 6.96 | 1599
BM170 | SW-20100810BM17-15 | 3 | 6.45 | 6.676 | 6.45 | 6.79 | 6.88 | 1399
BM170 | SW-20100810BM17-16 | 4 | 6.46 | 6.742 | 6.53 | 6.78 | 6.62 | 1199
BM170 | SW-20100810BM17-17 | 5 | 5.82 | 6.052 | 5.73 | 6.43 | 6.19 | 1024
BM170 | SW-20100810BM17-18 | 6 | 5.6 | 6.112 | 5.6 | 6.29 | 6.1 | 999
BM170 | SW-20100810BM17-19 | 7 | 5.4 | 5.146 | 5.05 | 5.54 | 5.27 | 799
BM170 | SW-20100810BM17-20 | 8 | 4.3 | 4.437 | 4.22 | 4.71 | 4.39 | 600
BM170 | SW-20100810BM17-21 | 9 | 3.98 | 3.995 | 3.87 | 4.35 | 3.98 | 400
BM170 | SW-20100810BM17-22 | 10 | 4.19 | 3.925 | 4.21 | 4.78 | 4.51 | 200
BM170 | SW-20100810BM17-23 | 11 | 6.15 | 5.301 | 6.19 | 6.41 | 6.38 | 2
BM171 | SW-20100810BM17-25 | 1 | 6.32 | 6.3 | 6.79 | 7.17 | 7 | 1689
BM171 | SW-20100810BM17-26 | 2 | 6.05 | 6.416 | 6.9 | 7.08 | 6.98 | 1599
BM171 | SW-20100810BM17-27 | 3 | 6.62 | 6.407 | 6.79 | 7.1 | 6.89 | 1399
BM171 | SW-20100810BM17-28 | 4 | 6.13 | 5.35 | 5.67 | 6.52 | 6.68 | 1199
BM171 | SW-20100810BM17-29 | 5 | 6.01 | 5.73 | 5.62 | 6.27 | 6.08 | 1000
BM171 | SW-20100810BM17-30 | 6 | 5.25 | 5.075 | 5.32 | 5.61 | 5.34 | 800
BM171 | SW-20100810BM17-31 | 7 | 4.25 | 4.284 | 4.19 | 4.75 | 4.39 | 600
BM171 | SW-20100810BM17-32 | 8 | 4.42 | 3.943 | 3.74 | 4.31 | 3.94 | 400
BM171 | SW-20100810BM17-33 | 9 | 4.15 | 4.023 | 4.24 | 4.76 | 4.48 | 200
BM171 | SW-20100810BM17-34 | 10 | 5.65 | 5.54 | 5.51 | 6.52 | 6.41 | 2
BM172 | SW-20100811BM17-1 | 1 | na | 6.112 | 6.18 | 6.78 | 6.63 | 1343
BM172 | SW-20100811BM17-2 | 2 | na | 5.677 | 6.05 | 6.54 | 6.36 | 1199
BM172 | SW-20100811BM17-3 | 3 | na | 5.591 | 6.06 | 6.32 | 6 | 1189
BM172 | SW-20100811BM17-4 | 4 | na | 5.042 | 5.63 | 6.16 | 5.84 | 1179
BM172 | SW-20100811BM17-5 | 5 | na | 4.997 | 6 | 6.19 | 6.02 | 1169
BM172 | SW-20100811BM17-6 | 6 | na | 4.813 | 5.67 | 6.06 | 5.88 | 1159
BM172 | SW-20100811BM17-7 | 7 | na | 4.46 | 5.58 | 5.98 | 5.79 | 1149
BM172 | SW-20100811BM17-8 | 8 | na | 4.87 | 5.79 | 5.97 | 5.73 | 1139
BM172 | SW-20100811BM17-9 | 9 | na | 4.655 | 5.36 | 5.91 | 5.71 | 1129
BM172 | SW-20100811BM17-10 | 10 | na | 4.824 | 5.37 | 5.91 | 5.73 | 1119
BM172 | SW-20100811BM17-11 | 11 | na | 4.787 | 5.81 | 5.93 | 5.75 | 1110
BM175 | SW-20100816BM18-1 | 2 | 5.44 |  | 5.23 | 5.79 | 5.63 | 1055
BM175 | SW-20100816BM18-2 | 4 | 5.23 |  | 5.26 | 5.84 | 5.56 | 999
BM175 | SW-20100816BM18-3 | 6 | 4.51 |  | 4.57 | 5.14 | 4.74 | 750
BM175 | SW-20100816BM18-4 | 8 | 3.96 |  | 3.84 | 4.3 | 3.93 | 500
BM175 | SW-20100816BM18-5 | 10 | 5.73 |  | 6.47 | 6.87 | 6.86 | 2
BM176 | SW-20100816BM18-7 | 2 | 6.14 |  | 6.24 | 6.53 | 6.68 | 1412
BM176 | SW-20100816BM18-8 | 4 | 5.78 |  | 5.88 | 6.28 | 5.95 | 1049
BM176 | SW-20100816BM18-9 | 6 | 5.35 |  | 4.63 | 5.13 | 4.76 | 750
BM176 | SW-20100816BM18-10 | 8 | 4.47 |  | 3.96 | 4.32 | 3.87 | 500
BM176 | SW-20100816BM18-11 | 10 | 6.52 |  | 6.19 | 6.62 | 6.52 | 2
BM177 | SW-20100816BM18-13 | 2 | 6.57 |  | 6.23 | 6.49 | 6.43 | 1325
BM177 | SW-20100816BM18-14 | 4 | 6 |  | 5.71 | 6.06 | 5.79 | 1099
BM177 | SW-20100816BM18-15 | 6 | 5.35 |  | 4.94 | 5.41 | 5.1 | 749
BM177 | SW-20100816BM18-16 | 8 | 4.22 |  | 3.78 | 4.35 | 3.96 | 500
BM177 | SW-20100816BM18-17 | 10 | 6.19 |  | 5.82 | 6.52 | 6.42 | 2
BM178 | SW-20100818BM18-1 | 2 | 6.81 |  | 6.85 | 7.3 | 7.05 | 1905
BM178 | SW-20100818BM18-2 | 4 | 6.62 |  | 6.65 | 6.89 | 6.66 | 1239
BM178 | SW-20100818BM18-3 | 6 | 6.42 |  | 6.41 | 6.71 | 6.49 | 1149
BM178 | SW-20100818BM18-4 | 8 | 4.3 |  | 3.86 | 4.6 | 4 | 500
BM178 | SW-20100818BM18-5 | 10 | 5.65 |  | 5.63 | 6.5 | 6.36 | 2
BM179 | SW-20100816BM18-9 | 2 | 6.61 |  | 6.48 | 6.91 | 6.78 | 1484
BM179 | SW-20100816BM18-10 | 4 | 6.32 |  | 5.94 | 6.68 | 6.45 | 1149
BM179 | SW-20100816BM18-11 | 6 | 5.71 |  | 5.85 | 6.32 | 6 | 1019
BM179 | SW-20100816BM18-12 | 8 | 4.02 |  | 3.87 | 4.61 | 4.04 | 500
BM179 | SW-20100816BM18-13 | 10 | 5.75 |  | 6.03 | 6.57 | 6.41 | 2
BM180 | SW-20100818BM18-15 | 2 | na |  | 5.26 | 6.01 | 5.91 | 1094
BM180 | SW-20100818BM18-16 | 4 | na |  | 4.91 | 5.95 | 5.7 | 999
BM180 | SW-20100818BM18-17 | 6 | na |  | 4.39 | 5.29 | 4.94 | 750
BM180 | SW-20100818BM18-18 | 8 | na |  | 3.69 | 4.41 | 3.99 | 500
BM180 | SW-20100818BM18-19 | 10 | na |  | 5.99 | 6.46 | 6.45 | 3
BM182 | SW-20100822BM19-1 | 2 | 5.77 |  | 5.66 | 5.97 | 5.79 | 1071 | 0.13
BM182 | SW-20100822BM19-2 | 4 | 5.57 |  | 5.08 | 5.79 | 5.61 | 1049 | 0.53
BM182 | SW-20100822BM19-3 | 6 | 5.3 |  | 5.28 | 5.7 | 5.47 | 999 | 0.19
BM182 | SW-20100822BM19-4 | 8 | 4.12 |  | 3.65 | 4.42 | 3.92 | 500 | 0.27
BM182 | SW-20100822BM19-5 | 10 | 5.5 |  | 6.2 | 6.6 | 6.5 | 2 | 0.3
BM183 | SW-20100822BM19-7 | 2 | 6.13 |  | 6.37 | 6.51 | 6.41 | 1240 | 0.04
BM183 | SW-20100822BM19-8 | 4 | 5.3 |  | 5.67 | 6.06 | 5.79 | 1049 | 0.12
BM183 | SW-20100822BM19-9 | 6 | 5.21 |  | 4.99 | 5.87 | 5.59 | 1000 | 0.6
BM183 | SW-20100822BM19-10 | 8 | 3.64 |  | 3.74 | 4.43 | 3.89 | 500 | 0.15
BM183 | SW-20100822BM19-11 | 10 | 5.35 |  | 6.38 | 6.64 | 6.68 | 2 | 0.3
BM184 | SW-20100822BM19-13 | 2 | 6.34 |  | 6.53 | 6.68 | 6.61 | 1338 | 0.08
BM184 | SW-20100822BM19-14 | 4 | 5.43 |  | 5.26 | 6.15 | 5.87 | 1119 | 0.61
BM184 | SW-20100822BM19-15 | 6 | 5.1 |  | 5.56 | 5.99 | 5.88 | 1079 | 0.32
BM184 | SW-20100822BM19-16 | 8 | 3.8 |  | 3.75 | 4.33 | 3.92 | 500 | 0.17
BM184 | SW-20100822BM19-17 | 10 | 5.74 |  | 6.64 | 6.87 | 6.99 | 2 | 0.35
BM185 | SW-20100822BM19-21 | 2 | 6.05 |  | 6.32 | 6.73 | 6.67 | 1478 | 0.35
BM185 | SW-20100822BM19-22 | 4 | 5.38 |  | 5.7 | 6.49 | 6.28 | 1179 | 0.58
BM185 | SW-20100822BM19-23 | 6 | 5.29 |  | 5.68 | 6.18 | 5.98 | 1049 | 0.3
BM185 | SW-20100822BM19-24 | 8 | 3.75 |  | 3.63 | 4.48 | 3.93 | 500 | 0.3
BM185 | SW-20100822BM19-25 | 10 | 6.32 |  | 6.79 | 7.02 | 7.37 | 2 | 0.58
BM186 | SW-20100823BM19-1 | 2 | 6.9 |  | 6.3 | 6.93 | 6.85 | 1641 | 0.55
BM186 | SW-20100823BM19-2 | 4 | 6.33 |  | 6.03 | 6.5 | 6.3 | 1150 | 0.27
BM186 | SW-20100823BM19-3 | 6 | 6.25 |  | 5.34 | 6.23 | 5.95 | 1049 | 0.61
BM186 | SW-20100823BM19-4 | 8 | 4.31 |  | 3.71 | 4.46 | 3.92 | 500 | 0.21
BM186 | SW-20100823BM19-5 | 10 | 6.21 |  | 6.31 | 6.53 | 6.72 | 2 | 0.41
BM187 | SW-20100823BM19-8 | 2 | 6.39 |  | 6.62 | 7.03 | 6.86 | 1688 | 0.24
BM187 | SW-20100823BM19-9 | 4 | 6.25 |  | 5.78 | 6.66 | 6.36 | 1149 | 0.58
BM187 | SW-20100823BM19-10 | 6 | 5.85 |  | 5.74 | 6.32 | 6.03 | 1049 | 0.29
BM187 | SW-20100823BM19-11 | 8 | 3.98 |  | 3.87 | 4.61 | 3.95 | 500 | 0.08
BM187 | SW-20100823BM19-12 | 10 | 6.15 |  | 6.45 | 6.6 | 6.87 | 3 | 0.42
BM188 | SW-20100823BM19-14 | 2 | 7 |  | 6.67 | 6.91 | 6.82 | 1588 | 0.15
BM188 | SW-20100823BM19-15 | 4 | 5.95 |  | 6.09 | 6.34 | 6.08 | 1119 | -0.01
BM188 | SW-20100823BM19-16 | 6 | 5.96 |  | 5.54 | 6.3 | 6.11 | 1089 | 0.57
BM188 | SW-20100823BM19-17 | 8 | 3.94 |  | 3.82 | 4.46 | 3.93 | 500 | 0.11
BM188 | SW-20100823BM19-18 | 10 | 6.01 |  | 6.75 | 7.07 | 7.33 | 3 | 0.58
BM189 | SW-20100823BM19-21 | 2 | 6.04 |  | 6.07 | 6.86 | 6.65 | 1377 | 0.58
BM189 | SW-20100823BM19-22 | 4 | 5.58 |  | 5.68 | 6.33 | 6.08 | 1149 | 0.4
BM189 | SW-20100823BM19-23 | 6 | 4.58 |  | 4.45 | 5.25 | 4.7 | 750 | 0.25
BM189 | SW-20100823BM19-24 | 8 | 3.8 |  | 3.73 | 4.37 | 3.91 | 500 | 0.18
BM189 | SW-20100823BM19-25 | 10 | 6.15 |  | 6.68 | 7.02 | 7.35 | 3 | 0.67
BM190 | SW-20100824BM19-1 | 2 | 6.46 |  | 6 | 6.6 | 6.49 | 1285 | 0.49
BM190 | SW-20100824BM19-2 | 4 | 6.47 |  | 5.65 | 6.21 | 6 | 1153 | 0.35
BM190 | SW-20100824BM19-3 | 6 | 6.4 |  | 5.61 | 6.15 | 5.92 | 1079 | 0.31
BM190 | SW-20100824BM19-4 | 8 | 4.14 |  | 3.68 | 4.35 | 3.85 | 500 | 0.17
BM190 | SW-20100824BM19-5 | 10 | 6.7 |  | 6 | 6.67 | 6.85 | 2 | 0.85
BM191 | SW-20100824BM19-7 | 2 | 6.42 |  | 6.16 | 6.9 | 6.78 | 1488 | 0.62
BM191 | SW-20100824BM19-8 | 4 | 5.71 |  | 5.16 | 6.1 | 5.7 | 1149 | 0.54
BM191 | SW-20100824BM19-9 | 6 | 6.71 |  | 5.77 | 6.29 | 5.99 | 1049 | 0.22
BM191 | SW-20100824BM19-10 | 8 | 3.93 |  | 3.58 | 4.38 | 3.92 | 500 | 0.34
BM191 | SW-20100824BM19-11 | 10 | 5.93 |  | 6.38 | 6.8 | 7.01 | 2 | 0.63
BM192 | SW-20100824BM19-15 | 2 | 6.69 |  | 6.81 | 7.17 | 6.95 | 1759 | 0.14
BM192 | SW-20100824BM19-17 | 4 | 6.46 |  | 5.69 | 6.61 | 6.36 | 1149 | 0.67
BM192 | SW-20100824BM19-18 | 6 | 5.77 |  | 5.76 | 6.28 | 6.02 | 1049 | 0.26
BM192 | SW-20100824BM19-19 | 8 | 4.06 |  | 3.63 | 4.48 | 4 | 500 | 0.37
BM192 | SW-20100824BM19-20 | 10 | 5.91 |  | 6.21 | 6.73 | 6.89 | 2 | 0.68
BM193 | SW-20100824BM19-22 | 2 | 6.59 |  | 6.42 | 7.1 | 6.95 | 1799 | 0.53
BM193 | SW-20100824BM19-23 | 4 | 6.53 |  | 6.23 | 6.69 | 6.48 | 1149 | 0.25
BM193 | SW-20100824BM19-24 | 6 | 6.24 |  | 5.79 | 6.41 | 6.17 | 1049 | 0.38
BM193 | SW-20100824BM19-25 | 8 | 4.15 |  | 3.86 | 4.39 | 4 | 500 | 0.14
BM193 | SW-20100824BM19-26 | 10 | 5.66 |  | 6.31 | 6.55 | 6.56 | 3 | 0.25
BM 194 | SW-20100830BM20-01 | 2 | 5.73 |  | 6.92 | na | 6.99 | 1890
BM 194 | SW-20100830BM20-02 | 4 | 5.7 |  | 6.22 | na | 6.51 | 1150 | 0.33525
BM 194 | SW-20100830BM20-03 | 6 | 5.85 |  | 5.98 | na | 4.69 | 1050
BM 194 | SW-20100830BM20-04 | 8 | 4.32 |  | 3.71 | na | 3.94 | 500
BM 194 | SW-20100830BM20-05 | 10 | 5.25 |  | 5.91 | na | 6.33 | 2
BM 195 | SW-20100830BM20-06 | 2 | 5.3 |  | 6.72 | na | 7.01 | 1890
BM 195 | SW-20100830BM20-07 | 4 | 6.7 |  | 6.13 | na | 6.46 | 1150
BM 195 | SW-20100830BM20-08 | 6 | 5.8 |  | 5.88 | na | 6.18 | 1050
BM 195 | SW-20100830BM20-09 | 8 | 4.8 |  | 3.8 | na | 3.89 | 500
BM 195 | SW-20100830BM20-10 | 10 | 5.6 |  | 5.98 | na | 6.35 | 2
BM 196 | SW-20100830BM20-11 | 2 | 11.2 |  | 6.84 | na | 7.02 | 1890
BM 196 | SW-20100830BM20-12 | 4 | 11.4 |  | 6.53 | na | 6.66 | 1150
BM 196 | SW-20100830BM20-13 | 6 | 9.2 |  | 6 | na | 6.23 | 1000
BM 196 | SW-20100830BM20-14 | 8 | 6.1 |  | 4.3 | na | 3.97 | 500
BM 196 | SW-20100830BM20-15 | 10 | 8.4 |  | 5.85 | na | 6.37 | 2
BM 197 | SW-20100831BM20-1 | 2 | 6.4 |  | 6.67 | na | 7.02 | 1890
BM 197 | SW-20100831BM20-2 | 4 | 5.9 |  | 6.57 | na | 6.4 | 1080
BM 197 | SW-20100831BM20-3 | 6 | 5.4 |  | 5.67 | na | 6.12 | 980
BM 197 | SW-20100831BM20-4 | 8 | 4 |  | 3.72 | na | 4.05 | 500
BM 197 | SW-20100831BM20-5 | 10 | 4.4 |  | 5.87 | na | 6.4 | 2
BM 198 | SW-20100831BM20-6 | 2 | 6.6 |  | 6.73 | na | 7.02 | 1890
BM 198 | SW-20100831BM20-7 | 4 | 6.7 |  | 5.97 | na | 6.29 | 1075
BM 198 | SW-20100831BM20-8 | 6 | 5.7 |  | 5.76 | na | 5.99 | 1010
BM 198 | SW-20100831BM20-9 | 8 | 4 |  | 3.87 | na | 4.05 | 500
BM 198 | SW-20100831BM20-10 | 10 | 4.9 |  | 6.02 | na | 6.44 | 2
BM 199 | SW-20100901BM20-01 | 2 | 9.1 |  | 6.37 | na | 6.7 | 1280
BM 199 | SW-20100901BM20-02 | 4 | 6.2 |  | 5.68 | na | 5.9 | 1060
BM 199 | SW-20100901BM20-03 | 6 | 6 |  | 5.64 | na | 5.8 | 1000
BM 199 | SW-20100901BM20-04 | 8 | 4.1 |  | 3.85 | na | 4 | 500
BM 199 | SW-20100901BM20-05 | 10 | 5.7 |  | 6.08 | na | 6.5 | 2
BM 200 | SW-20100901BM20-06 | 2 | 7 |  | 6.03 | na | 6.2 | 1135
BM 200 | SW-20100901BM20-07 | 4 | 6.5 |  | 5.57 | na | 5.8 | 1060
BM 200 | SW-20100901BM20-08 | 6 | 7 |  | 5.63 | na | 5.8 | 1000
BM 200 | SW-20100901BM20-09 | 8 | 5.4 |  | 3.83 | na | 4.1 | 500
BM 200 | SW-20100901BM20-10 | 10 | 5.8 |  | 6.12 | na | 6.5 | 2
BM 201 | SW-20100901BM20-11 | 2 | 8.7 |  | 6.01 | na | 6.2 | 1140
BM 201 | SW-20100901BM20-12 | 4 | 7.8 |  | 5.29 | na | 5.9 | 1030
BM 201 | SW-20100901BM20-13 | 6 | 6.5 |  | 4.92 | na | 5.2 | 850
BM 201 | SW-20100901BM20-14 | 8 | 5.1 |  | 3.91 | na | 3.9 | 500
BM 201 | SW-20100901BM20-15 | 10 | 6 |  | 6.05 | na | 6.5 | 2
BM 202 | SW-20100901BM20-16 | 2 | 6.1 |  | 5.59 | na | 5.8 | 962
BM 202 | SW-20100901BM20-17 | 4 | 6.7 |  | 5.02 | na | 5.7 | 900
BM 202 | SW-20100901BM20-18 | 6 | 5.7 |  | 4.26 | na | 4.6 | 700
BM 202 | SW-20100901BM20-19 | 8 | 5 |  | 3.84 | na | 4 | 500
BM 202 | SW-20100901BM20-20 | 10 | 6.5 |  | 6.07 | na | 6.5 | 2
BM 203 | SW-20100903BM21-1 | 2 | N/A |  | 5.86 | 6.5 | 6.1 | 1119
BM 203 | SW-20100903BM21-2 | 4 | N/A |  | 5.65 | 6.3 | 5.9 | 1050
BM 203 | SW-20100903BM21-3 | 6 | N/A |  | 4.21 | 4.9 | 4.3 | 641
BM 203 | SW-20100903BM21-4 | 8 | N/A |  | 3.82 | 4.5 | 4 | 500
BM 203 | SW-20100903BM21-5 | 10 | N/A |  | 6.06 | 6.8 | 6.4 | 3
BM 204 | SW-20100903BM21-6 | 2 | N/A |  | 6.23 | 6.7 | 6.5 | 1228
BM 204 | SW-20100903BM21-7 | 4 | N/A |  | 5.61 | 6.3 | 5.9 | 1049
BM 204 | SW-20100903BM21-8 | 6 | N/A |  | 4.86 | 5.1 | 5 | 800
BM 204 | SW-20100903BM21-9 | 8 | N/A |  | 3.86 | 4.4 | 4 | 500
BM 204 | SW-20100903BM21-10 | 10 | N/A |  | 6.07 | 6.7 | 6.4 | 2
BM 205 | SW20100904BM21-01 | 2 | 6.5 |  | 6.46 | 6.8 | 6.7 | 1396
BM 205 | SW20100904BM21-02 | 4 | 8.2 |  | 6.1 | 6.7 | 6.3 | 1150
BM 205 | SW20100904BM21-03 | 6 | 6.5 |  | 5.82 | 6.4 | 6 | 1079
BM 205 | SW20100904BM21-04 | 8 | 4.6 |  | 3.86 | 4.4 | 3.9 | 500
BM 205 | SW20100904BM21-05 | 10 | 5.6 |  | 5.99 | 6.5 | 6.3 | 3
BM 206 | SW20100904BM21-06 | 2 | 7.3 |  | 6.56 | 7.1 | 6.9 | 1838
BM 206 | SW20100904BM21-07 | 4 | 7.8 |  | 6.14 | 6.8 | 6.4 | 1199
BM 206 | SW20100904BM21-08 | 6 | 6.2 |  | 5.72 | 6.4 | 5.9 | 1049
BM 206 | SW20100904BM21-09 | 8 | 6.1 |  | 3.74 | 4.5 | 3.9 | 500
BM 206 | SW20100904BM21-10 | 10 | 8.6 |  | 5.99 | 6.6 | 2.8 | 3
BM 207 | SW20100904BM21-11 | 2 | 7.1 |  | 6.18 | 6.5 | 6.8 | 1477
BM 207 | SW20100904BM21-12 | 4 | 5.8 |  | 5.57 | 4.4 | 5.9 | 1156
BM 207 | SW20100904BM21-13 | 6 | 7.2 |  | 5.43 | 6.2 | 5.9 | 1050
BM 207 | SW20100904BM21-14 | 8 | 7.3 |  | 3.61 | 6.2 | 3.9 | 500
BM 207 | SW20100904BM21-15 | 10 | 7.7 |  | 5.77 | 7.1 | 6.3 | 3
BM 208 | SW20100904BM21-16 | 2 | 8.5 |  | 6.47 | 7.2 | 7 | 1888
BM 208 | SW20100904BM21-17 | 4 | 7.4 |  | 6.56 | 6.5 | 7 | 1200
BM 208 | SW20100904BM21-18 | 6 | 7.4 |  | 5.4 | 6.2 | 5.8 | 1049
BM 208 | SW20100904BM21-19 | 8 | 5.3 |  | 3.65 | 4.1 | 3.9 | 500
BM 208 | SW20100904BM21-20 | 10 | 6.1 |  | 5.76 | 6.6 | 6.4 | 3
BM 209 | SW20100905BM21-01 | 2 | NA |  | 2.49 | 7.3 | 6.7 | 1848
BM 209 | SW20100905BM21-02 | 4 | NA |  | 4.95 | 7.2 | 6.5 | 1199
BM 209 | SW20100905BM21-03 | 6 | NA |  | 5.64 | 6.4 | 6.1 | 1059
BM 209 | SW20100905BM21-04 | 7 | NA |  | 3.71 | 4.5 | 3.9 | 500
BM 209 | SW20100905BM21-05 | 10 | NA |  | 5.83 | 6.6 | 6.4 | 2
BM 210 | SW20100905BM21-06 | 2 | 7 |  | 6.52 | 7.5 | 7 | 1888
BM 210 | SW20100905BM21-07 | 4 | 6 |  | 5.56 | 6.4 | 6 | 1139
BM 210 | SW20100905BM21-08 | 6 | 5.7 |  | 5.52 | 6.3 | 5.7 | 1050
BM 210 | SW20100905BM21-09 | 8 | 3.8 |  | 3.59 | 4.3 | 3.8 | 500
BM 210 | SW20100905BM21-10 | 10 | 6.4 |  | 5.81 | 6.5 | 6.4 | 2
BM 211 | SW20100905BM21-11 | 2 | NA |  | N/A | 7.3 | 6.5 | 1889
BM 211 | SW20100905BM21-12 | 4 | NA |  | N/A | 6.9 | 6.5 | 1199
BM 211 | SW20100905BM21-13 | 6 | NA |  | N/A | 6.5 | 6.1 | 1059
BM 211 | SW20100905BM21-14 | 8 | NA |  | N/A | 4.3 | 3.8 | 500
BM 211 | SW20100905BM21-15 | 10 | NA |  | N/A | 6.7 | 6.4 | 2
StationID | SampleID | Niskin Bottle# | DO_Membrane | DO_Winkler | DO_Hach | DO_Optical | DO_Seabird | Depth
|  |  | (mg/L) | (mg/L) | (mg/L) | (mg/L) | (mg/L) | Depth [m]
|  | Data set from Brooks McCall Cruise 15 (13 sites, 65 stations) | [5 empty] | Total data set from Brooks McCall (26 sites, 158 stations)
|  | Correlation | Seabird | Hach | Optical | Membrane |  | Correlation | Seabird | Hach | Optical | Membrane
|  | Winkler | - | - | - | - |  | Winkler | 0.86104 | 0.86749 | 0.85121 | 0.87223
|  | Seabird |  | 0.96932 | 0.90658 | 0.79454 |  | Seabird |  | 0.96236 | 0.89072 | 0.81001
|  | Hach |  |  | 0.89774 | 0.8427 |  | Hach |  |  | 0.85403 | 0.78075
|  | Optical |  |  |  | 0.81681 |  | Optical |  |  |  | 0.88618
This total table is now outdated
Data set from Brooks McCall Cruise 16 (13 sites, 50 stations)
|  | Correlation | Seabird | Hach | Optical | Membrane
|  | Winkler | 0.82667 | 0.85931 | 0.80965 | 0.80779
|  | Seabird |  | 0.96419 | 0.98093 | 0.89426
|  | Hach |  |  | 0.95841 | 0.94736
|  | Optical |  |  |  | 0.90104
Data set from Brooks McCall Cruise 17 (4 sites, 43 stations)
|  | Correlation | Seabird | Hach | Optical | Membrane
|  | Winkler | 0.86587 | 0.85432 | 0.88158 | 0.90178
|  | Seabird |  | 0.95792 | 0.99422 | 0.93121
|  | Hach |  |  | 0.96572 | 0.92905
|  | Optical |  |  |  | 0.9326
Data set from Brooks McCall Cruise 18 (6 sites, 30 stations)
|  | Correlation | Seabird | Hach | Optical | Membrane
|  | Winkler | - | - | - | -
|  | Seabird |  | 0.98505 | 0.99437 | 0.93748
|  | Hach |  |  | 0.98094 | 0.9536
|  | Optical |  |  |  | 0.94005
Data set from Brooks McCall Cruise 19 (12 sites, 60 stations)
|  | Correlation | Seabird | Hach | Optical | Membrane
|  | Winkler | - | - | - | -
|  | Seabird |  | 0.98505 | 0.99058 | 0.90258
|  | Hach |  |  | 0.97991 | 0.89538
|  | Optical |  |  |  | 0.92393
Data set from Brooks McCall Cruise 20 (12 sites, 60 stations)
|  | Correlation | Seabird | Hach | Optical | Membrane
|  | Winkler | - | - | - | -
|  | Seabird |  | 0.90678 |  | 0.59714
|  | Hach |  |  |  | 5.58
Optical
Data set from Brooks McCall Cruise 21 (12 sites, 60 stations)
|  | Correlation | Seabird | Hach | Optical | Membrane
|  | Winkler | - | - | - | -
|  | Seabird |  | 0.77382 | 0.59902 | 0.22991 | Calibration was unavailable for stations BM203, BM204
|  | Hach |  |  |  | 0.52587
Optical
Note on D296: Owner: Amorales solution for calibration unavailable
Note on D306: Owner: Amorales Titrated solution provided for calibration
Note on D326: Owner: Equipment failure no display
Note on C329: Owner: Amorales Bottle 7 used instead of 8 water is from same depth

## OV Data
StationID | SampleID | Niskin Bottle# | DO_Membrane | DO_Winkler | DO_Hach | DO_Optical | DO_Seabird | Depth | Data to the right are sorted for paired comparisons | StationID | SampleID | Niskin Bottle# | DO_Membrane | DO_Winkler | DO_Hach | DO_Optical | DO_Seabird | Depth
|  |  | (mg/L) | (mg/L) | (mg/L) | (mg/L) | (mg/L) | Depth [m] |  |  |  |  | (mg/L) | (mg/L) | (mg/L) | (mg/L) | (mg/L) | Depth [m]
OV149 | SW-20100801-OV12-001 | 1 | 5.33 | 5.71 |  | 5.68 | 5.76 | 1149 |  | OV153 | SW-20100802-OV12-001 | 1 | 6.25 | 6.79 |  | 6.61 | 6.78 | 1450
OV149 | SW-20100801-OV12-002 | 2 | 5.23 | 5.97 |  | 6.04 | 5.81 | 1010 |  | OV152 | SW-20100801-OV12-032 | 1 | 6.12 | 5.87 |  | 6.78 | 6.5 | 1350
OV149 | SW-20100801-OV12-007 | 5 | 3.38 | 3.88 |  | 4.2 | 3.91 | 400 |  | OV152 | SW-20100801-OV12-038 | 7 | 6.03 | 6.8 |  | 7.64 | 7.29 | 50
OV149 | SW-20100801-OV12-010 | 7 | 5.48 | 7.2 |  | 7.56 | 7.34 | 50 |  | OV153 | SW-20100802-OV12-004 | 2 | 5.86 | 5.72 |  | 6.67 | 6.39 | 1300
OV150 | SW-20100801-OV12-013 | 1 | 4.57 | 4.9 |  | 5.03 | 4.94 | 750 |  | OV152 | SW-20100801-OV12-034 | 3 | 5.84 | 6.46 |  | 6.99 | 6.49 | 1250
OV150 | SW-20100801-OV12-014 | 2 | 3.58 | 4.36 |  | 4.62 | 4.39 | 650 |  | OV152 | SW-20100801-OV12-037 | 6 | 5.73 | 7.45 |  | 7.73 | 3.98 | 400
OV150 | SW-20100801-OV12-015 | 3 | 3.42 | 3.93 |  | 4.29 | 3.97 | 500 |  | OV152 | SW-20100801-OV12-033 | 2 | 5.73 | 6.06 |  | 6.76 | 6.51 | 1300
OV150 | SW-20100801-OV12-016 | 4 | 3.39 | 4.13 |  | 4.65 | 4.19 | 300 |  | OV154 | SW-20100802-OV12-014 | 2 | 5.64 | 6.31 |  | 6.82 | 6.29 | 1272
OV150 | SW-20100801-OV12-017 | 5 | 3.25 | 3.46 |  | 4.37 | 4.03 | 233 |  | OV154 | SW-20100802-OV12-013 | 1 | 5.63 | 6.07 |  | 7.02 | 6.63 | 1350
OV150 | SW-20100801-OV12-018 | 6 | 3.81 | 4.89 |  | 5.3 | 4.97 | 128 |  | OV152 | SW-20100801-OV12-035 | 4 | 5.59 | 6.26 |  | 6.71 | 5.92 | 1192
OV150 | SW-20100801-OV12-019 | 7 | 5.51 | 6.08 |  | 7.45 | 7.24 | 50 |  | OV150 | SW-20100801-OV12-019 | 7 | 5.51 | 6.08 |  | 7.45 | 7.24 | 50
OV150 | SW-20100801-OV12-020 | 8 | 4.83 | 5.32 |  | 6.75 | 6.44 | 3 |  | OV149 | SW-20100801-OV12-010 | 7 | 5.48 | 7.2 |  | 7.56 | 7.34 | 50
OV151 | SW-20100801-OV12-021 | 1 | 5.22 | 5.73 |  | 5.77 | 5.77 | 1060 |  | OV153 | SW-20100802-OV12-005 | 3 | 5.44 | 5.87 |  | 6.3 | 5.69 | 1206
OV151 | SW-20100801-OV12-022 | 2 | 4.91 | 5.63 |  | 5.89 | 5.68 | 1020 |  | OV152 | SW-20100801-OV12-036 | 5 | 5.38 | 6.1 |  | 6.31 | 5.83 | 1165
OV151 | SW-20100801-OV12-023 | 3 | 3.52 | 4.22 |  | 4.56 | 4.2 | 600 |  | OV153 | SW-20100802-OV12-010 | 7 | 5.35 | 6.87 |  | 7.4 | 7.17 | 50
OV151 | SW-20100801-OV12-024 | 4 | 3.53 | 3.38 |  | 4.28 | 3.92 | 351 |  | OV154 | SW-20100802-OV12-016 | 4 | 5.34 | 5.84 |  | 6.42 | 5.59 | 1169
OV151 | SW-20100801-OV12-025 | 5 | 3.46 | 4.21 |  | 4.57 | 4.24 | 300 |  | OV149 | SW-20100801-OV12-001 | 1 | 5.33 | 5.71 |  | 5.68 | 5.76 | 1149
OV151 | SW-20100801-OV12-026 | 6 | 3.91 | 4.83 |  | 5.24 | 4.89 | 150 |  | OV154 | SW-20100802-OV12-019 | 7 | 5.26 | 6.73 |  | 7.34 | 7.11 | 50
OV151 | SW-20100801-OV12-029 | 7 | 5.14 | 7.17 |  | 7.61 | 7.43 | 50 |  | OV149 | SW-20100801-OV12-002 | 2 | 5.23 | 5.97 |  | 6.04 | 5.81 | 1010
OV151 | SW-20100801-OV12-031 | 8 | 4.29 | 6.35 |  | 6.95 | 6.51 | 3 |  | OV151 | SW-20100801-OV12-021 | 1 | 5.22 | 5.73 |  | 5.77 | 5.77 | 1060
OV152 | SW-20100801-OV12-032 | 1 | 6.12 | 5.87 |  | 6.78 | 6.5 | 1350 |  | OV155 | SW-20100802-OV12-022 | 2 | 5.18 | 5.83 |  | 6.22 | 5.84 | 1020
OV152 | SW-20100801-OV12-033 | 2 | 5.73 | 6.06 |  | 6.76 | 6.51 | 1300 |  | OV154 | SW-20100802-OV12-015 | 3 | 5.15 | 5.67 |  | 6.38 | 5.92 | 1205
OV152 | SW-20100801-OV12-034 | 3 | 5.84 | 6.46 |  | 6.99 | 6.49 | 1250 |  | OV151 | SW-20100801-OV12-029 | 7 | 5.14 | 7.17 |  | 7.61 | 7.43 | 50
OV152 | SW-20100801-OV12-035 | 4 | 5.59 | 6.26 |  | 6.71 | 5.92 | 1192 |  | OV155 | SW-20100802-OV12-021 | 1 | 5.05 | 5.87 |  | 6.08 | 5.91 | 1060
OV152 | SW-20100801-OV12-036 | 5 | 5.38 | 6.1 |  | 6.31 | 5.83 | 1165 |  | OV151 | SW-20100801-OV12-022 | 2 | 4.91 | 5.63 |  | 5.89 | 5.68 | 1020
OV152 | SW-20100801-OV12-037 | 6 | 5.73 | 7.45 |  | 7.73 | 3.98 | 400 |  | OV150 | SW-20100801-OV12-020 | 8 | 4.83 | 5.32 |  | 6.75 | 6.44 | 3
OV152 | SW-20100801-OV12-038 | 7 | 6.03 | 6.8 |  | 7.64 | 7.29 | 50 |  | OV153 | SW-20100802-OV12-007 | 5 | 4.61 | 5.36 |  | 5.89 | 5.75 | 1050
OV152 | SW-20100801-OV12-039 | 8 | 4.38 | 4.66 |  | 6.89 | 6.45 | 3 |  | OV153 | SW-20100802-OV12-011 | 8 | 4.58 | 6.35 |  | 6.89 | 6.49 | 3
OV153 | SW-20100802-OV12-001 | 1 | 6.25 | 6.79 |  | 6.61 | 6.78 | 1450 |  | OV150 | SW-20100801-OV12-013 | 1 | 4.57 | 4.9 |  | 5.03 | 4.94 | 750
OV153 | SW-20100802-OV12-004 | 2 | 5.86 | 5.72 |  | 6.67 | 6.39 | 1300 |  | OV152 | SW-20100801-OV12-039 | 8 | 4.38 | 4.66 |  | 6.89 | 6.45 | 3
OV153 | SW-20100802-OV12-005 | 3 | 5.44 | 5.87 |  | 6.3 | 5.69 | 1206 |  | OV151 | SW-20100801-OV12-031 | 8 | 4.29 | 6.35 |  | 6.95 | 6.51 | 3
OV153 | SW-20100802-OV12-012 | 4 |  | 5.83 |  |  | 5.72 | 1157 |  | OV154 | SW-20100802-OV12-020 | 8 | 4.22 | 6.32 |  | 6.86 | 6.48 | 3
OV153 | SW-20100802-OV12-007 | 5 | 4.61 | 5.36 |  | 5.89 | 5.75 | 1050 |  | OV151 | SW-20100801-OV12-026 | 6 | 3.91 | 4.83 |  | 5.24 | 4.89 | 150
OV153 | SW-20100802-OV12-009 | 6 |  | 3.81 |  |  | 3.78 | 400 |  | OV150 | SW-20100801-OV12-018 | 6 | 3.81 | 4.89 |  | 5.3 | 4.97 | 128
OV153 | SW-20100802-OV12-010 | 7 | 5.35 | 6.87 |  | 7.4 | 7.17 | 50 |  | OV150 | SW-20100801-OV12-014 | 2 | 3.58 | 4.36 |  | 4.62 | 4.39 | 650
OV153 | SW-20100802-OV12-011 | 8 | 4.58 | 6.35 |  | 6.89 | 6.49 | 3 |  | OV151 | SW-20100801-OV12-024 | 4 | 3.53 | 3.38 |  | 4.28 | 3.92 | 351
OV154 | SW-20100802-OV12-013 | 1 | 5.63 | 6.07 |  | 7.02 | 6.63 | 1350 |  | OV151 | SW-20100801-OV12-023 | 3 | 3.52 | 4.22 |  | 4.56 | 4.2 | 600
OV154 | SW-20100802-OV12-014 | 2 | 5.64 | 6.31 |  | 6.82 | 6.29 | 1272 |  | OV151 | SW-20100801-OV12-025 | 5 | 3.46 | 4.21 |  | 4.57 | 4.24 | 300
OV154 | SW-20100802-OV12-015 | 3 | 5.15 | 5.67 |  | 6.38 | 5.92 | 1205 |  | OV150 | SW-20100801-OV12-015 | 3 | 3.42 | 3.93 |  | 4.29 | 3.97 | 500
OV154 | SW-20100802-OV12-016 | 4 | 5.34 | 5.84 |  | 6.42 | 5.59 | 1169 |  | OV150 | SW-20100801-OV12-016 | 4 | 3.39 | 4.13 |  | 4.65 | 4.19 | 300
OV154 | SW-20100802-OV12-018 | 6 | 3.15 | 3.8 |  | 4.26 | 3.84 | 400 |  | OV149 | SW-20100801-OV12-007 | 5 | 3.38 | 3.88 |  | 4.2 | 3.91 | 400
OV154 | SW-20100802-OV12-019 | 7 | 5.26 | 6.73 |  | 7.34 | 7.11 | 50 |  | OV150 | SW-20100801-OV12-017 | 5 | 3.25 | 3.46 |  | 4.37 | 4.03 | 233
OV154 | SW-20100802-OV12-020 | 8 | 4.22 | 6.32 |  | 6.86 | 6.48 | 3 |  | OV154 | SW-20100802-OV12-018 | 6 | 3.15 | 3.8 |  | 4.26 | 3.84 | 400
OV155 | SW-20100802-OV12-021 | 1 | 5.05 | 5.87 |  | 6.08 | 5.91 | 1060 |  | OV157 | SW-20100802-OV12-048 | 3 |  | 6.79 |  | 7.36 | 6.83 | 1500
OV155 | SW-20100802-OV12-022 | 2 | 5.18 | 5.83 |  | 6.22 | 5.84 | 1020 |  | OV157 | SW-20100802-OV12-049 | 4 |  | 6.87 |  | 7.25 | 6.77 | 1400
OV156 | SW-20100802-OV12-032 | 1 |  | 6.69 |  | 6.26 | 6.75 | 1340 |  | OV157 | SW-20100802-OV12-050 | 5 |  | 6.62 |  | 7.17 | 6.66 | 1300
OV156 | SW-20100802-OV12-033 | 2 |  | 6.54 |  | 6.66 | 6.56 | 1250 |  | OV156 | SW-20100802-OV12-034 | 3 |  | 6.35 |  | 6.7 | 6.41 | 1200
OV156 | SW-20100802-OV12-034 | 3 |  | 6.35 |  | 6.7 | 6.41 | 1200 |  | OV156 | SW-20100802-OV12-033 | 2 |  | 6.54 |  | 6.66 | 6.56 | 1250
OV156 | SW-20100802-OV12-035 | 4 |  | 5.33 |  | 5.76 | 6.16 | 1150 |  | OV157 | SW-20100802-OV12-051 | 6 |  | 6.09 |  | 6.65 | 6.43 | 1200
OV156 | SW-20100802-OV12-036 | 5 |  | 5.84 |  | 6.33 | 4.73 | 1131 |  | OV157 | SW-20100802-OV12-053 | 8 |  | 5.86 |  | 6.42 | 5.86 | 1069
OV156 | SW-20100802-OV12-039 | 6 |  | 4.55 |  | 5.01 | 4.63 | 1124 |  | OV156 | SW-20100802-OV12-036 | 5 |  | 5.84 |  | 6.33 | 4.73 | 1131
OV156 | SW-20100802-OV12-040 | 7 |  | 5.13 |  | 5.52 | 5.1 | 1100 |  | OV156 | SW-20100802-OV12-042 | 9 |  | 5.81 |  | 6.29 | 5.83 | 1025
OV156 | SW-20100802-OV12-041 | 8 |  | 5.36 |  | 5.84 | 5.38 | 1085 |  | OV156 | SW-20100802-OV12-032 | 1 |  | 6.69 |  | 6.26 | 6.75 | 1340
OV156 | SW-20100802-OV12-042 | 9 |  | 5.81 |  | 6.29 | 5.83 | 1025 |  | OV157 | SW-20100802-OV12-052 | 7 |  | 5.61 |  | 6.18 | 5.46 | 1132
OV157 | SW-20100802-OV12-048 | 3 |  | 6.79 |  | 7.36 | 6.83 | 1500 |  | OV156 | SW-20100802-OV12-041 | 8 |  | 5.36 |  | 5.84 | 5.38 | 1085
OV157 | SW-20100802-OV12-049 | 4 |  | 6.87 |  | 7.25 | 6.77 | 1400 |  | OV156 | SW-20100802-OV12-035 | 4 |  | 5.33 |  | 5.76 | 6.16 | 1150
OV157 | SW-20100802-OV12-050 | 5 |  | 6.62 |  | 7.17 | 6.66 | 1300 |  | OV156 | SW-20100802-OV12-040 | 7 |  | 5.13 |  | 5.52 | 5.1 | 1100
OV157 | SW-20100802-OV12-051 | 6 |  | 6.09 |  | 6.65 | 6.43 | 1200 |  | OV156 | SW-20100802-OV12-039 | 6 |  | 4.55 |  | 5.01 | 4.63 | 1124
OV157 | SW-20100802-OV12-052 | 7 |  | 5.61 |  | 6.18 | 5.46 | 1132 |  | OV153 | SW-20100802-OV12-012 | 4 |  | 5.83 |  |  | 5.72 | 1157
OV157 | SW-20100802-OV12-053 | 8 |  | 5.86 |  | 6.42 | 5.86 | 1069 |  | OV153 | SW-20100802-OV12-009 | 6 |  | 3.81 |  |  | 3.78 | 400
Data set from Ocean Veritas Cruise 12 (9 stations, 60 stations)
|  | Correlation | Seabird | Optical | Membrane
|  | Winkler | 0.81797 | 0.92353 | 0.84483
|  | Seabird |  | 0.86217 | 0.75225
|  | Optical |  |  | 0.82106